SAFETY PRECAUTIONS FOR MOCVD SYSTEM

Do not operate an MOCVD system if you are alone in room 115.  We require at least 2 people in room 115 while operating the system.

The first and foremost area where one has to be careful while operating the system is while Turning on/Turning off the system.

Procedure for turning on the MOCVD system

Step 1 (Very important for everyone’s safety):

1.Check and make sure that exhaust fans are on and operating.  If they are not on, do not operate MOCVD until they have been returned to operation.

Step 2 (power on the MOCVD system):

1.Turn on the chiller placed behind the MOCVD system.

2.Turn on the power supply for the RF generator.

3.Power on the main controller of the MOCVD by pushing two buttons  on the right middle of the main control panel and by switching in three bar-switches on the right bottom of the main control  panel.

Step 3(pump Hydrogen purifier):

1.Turn on the small pump by pushing the S-pump button placed on the glove box.

2.Open the red valve(connecting the tube from MO sources to the small pump)located at the front of the small pump

3.Open the green valve located at near Hydrogen cylinder(behind the MOCVD).

Whenever this green valve open hydrogen cylinder has to be closing.

4.open two green valves located at up behind hydrogen purifier

5.Wait for 2 min to see the temperature increase from the front panel of hydrogen purifier

6.close two green valves in #4.

7.close the green valve in # 3 and then close the red valve in #2

8.Turn off the small pump by pushing the S-pump button.

Step 4 :

1.open the 5.0 grade Nitrogen cylinder and open the screw valve by rotating ,finally adjusting outlet pressure

2.Repeat #1 for 5.0 Hydrogen and for normal nitrogen.

Step 4(Hydrogen purifier)

1.Open two red valves located at each end of oxygen trap.

2.open the green valve located at the next to the red valve after oxygen trap.

3.open the screw valve by rotating located on the front of the hydrogen purifier.

4.open the green valve located at down behind hydrogen purifier

5.set the flow rate of breed gas to around 0.5l/min.The setting valve is on the right down of the hydrogen purifier front side.

EMERGENCY PROCEDURES:

IF HYDROGEN ALARM(ON PURIFIER) SOUNDS

- CLOSE TMGa(if open)

a.Turn off Valve H1 on switch board-if in computer program switch to manual first.

b.Close valves on TMGa source.

c.Turn off valve H2 on switch board

d.Same procedure for any MO Source open

2. Switch to Nitrogen flow(if hydrogen was flowing)

a. Turn off valve A1 and open Valve A3

3.Turn off RF generator

4.Shut down Hydrogen

a.Close ultra pure hydrogen valve and bleed valve on Purifier

b.Close all Hydrogen feed valves(hand operated)

c.Close valve on Hydrogen Tank.

In case of severe Emergency call 911.

GENERAL PRECAUTIONS:

1.Never leave the machine when it is running.

2.Always check to make sure the exhaust fans are operating.

3.Always have an eye on the temperature(during high temperature growth) and on the system pressure.

4.In case of any abnormalities resort to emergency shut down.

5.Never touch the RF coil when the MOCVD  is operating.

6.Always wear a robe and lab footwear and goggles  while inside the MOCVD room and doing growth.

7.If you have a doubt on the operating procedures do not hesistate to ask the seniors.

8.The MOCVD ROOM IS SUPPOSED TO BE KEPT CLEAN OF DUST ALL THE TIME.Lock the door whenever you use it.

Procedures for handling the MO Sources:

1.Always use the MO sources in accordance with the Material Safety Data sheet.So go through it before you start work.

2.The MO sources are spontaneously combustible so pack them carefull after use and store them in the big yellow cabinets provided for that purpose.

3.Keep the MO sources under inert atmosphere and keep their containers dry always as it is dangerous when they are wet.

TRAINING  RULES FOR NEW STUDENTS IN MOCVD GROWTH

1.The training phase for a new student in the MOCVD growth typically last for about an year wherein the new student has to spend time observing a senior operate the system.The new student is required to make a note of all the controls/processes involved in the lab notebook provided.

2.Always read the Material Safety and Data Sheets before you enter the lab to watch/learn how to operate the MOCVD system.

3.When you are new, never touch or try to operate anything unless you have the go ahead from our Advisors and the person who has trained you.

4.During the training period,make sure you operate any equipment only under an experienced person’s supervision.NEVER operate it even though YOU feel comfortable doing it.

5.Understand the safety precautions to be followed while dealing with the M.O.sources and the other gases that we use in the MOCVD room.

6.Be very careful when you are turning on/off any of the valves when you are new because a fraction of second in negligence and you will be putting yourself and others around you in risk.

7.Make a note of the procedures involved in operating the MOCVD in the lab note book provided to you.This should serve as a reference manual too when you start running the machine independently.

8.Make a sketch of the MOCVD system and the gas flow lines in the lab note book so that you will understand the flow better and the important valves/lines to watch out for.

9.Always be attentive when you are in the room and do not take your eyes off the system pressure/flow meters.

10.Follow proper dress code and when in doubt always consult with the seniors.